TECHNICAL DOCUMENTATION

BTS Brewer #66639

Brewer Transportable Spectrometer — Kipp & Zonen

Total ozone column · UV spectrophotometry · Transportable reference

FMI — Sodankylä Arctic Research Station  ·  67°22’N, 26°38’E

# 1. Instrument Overview

The BTS (Brewer Transportable Spectrometer) #66639 is a portable version of the standard Kipp & Zonen Brewer spectrophotometer, designed to be deployed at field sites for campaign measurements and inter-comparisons. At Sodankylä, the BTS operates alongside the fixed Brewer MkIII to provide an independent measurement of the total ozone column (TOC) and serves as a cross-validation reference within the local instrument network.

Like all Brewer instruments, the BTS measures direct solar UV irradiance at five discrete wavelengths in the Hartley–Huggins ozone absorption band (306–320 nm). The differential absorption between wavelengths yields the ozone column in Dobson Units (DU). Its transportable design makes it particularly useful for NDACC inter-comparison campaigns and as a travelling reference to verify the calibration of fixed network instruments.

Note: The BTS Brewer Transportable Spectrometer (#66639) is a distinct instrument from the BTS DIAL Lidar also operated at Sodankylä. The DIAL lidar measures vertical ozone profiles by laser backscatter. The BTS described here measures total column ozone from direct solar UV photometry, and its data are stored as daily CSV files in ./BTS_data/.

Key specifications Instrument: Brewer Transportable Spectrometer (BTS) Serial number: #66639 Manufacturer: Kipp & Zonen (Netherlands) Measurement wavelengths: 306.3, 310.1, 313.5, 316.8, 320.1 nm Spectral resolution: ~0.6 nm (FWHM) Measurement geometry: Direct sun Derived product: Total ozone column (DU), Ozone uncertainty (DU) Temporal resolution: ~2–5 min (continuous during daylight) Accuracy (O₃): <1% (after WMO calibration) Data storage: Daily CSV files in ./BTS_data/

# 2. Measurement Principle

## 2.1  UV direct-sun photometry

The BTS uses the same measurement principle as all Brewer instruments. It measures the direct solar irradiance at five UV wavelengths and computes a weighted log-ratio (the F-ratio) that is sensitive to ozone absorption but cancels contributions from Rayleigh scattering, SO₂, and instrumental drift. The ozone column Ω is derived from Beer–Lambert’s law:

Ω = [F_meas − F_0] / (β × μ)

where F_0 is the extraterrestrial constant (determined by Langley calibration), β is the differential ozone absorption coefficient, and μ = 1/cos(SZA) is the air mass factor.

## 2.2  Transportable use and calibration

The BTS is designed to be quickly deployed at field sites while maintaining traceability to the WMO World Standard Group of reference Brewers. In practice, the BTS is calibrated before and after each deployment campaign, and its measurements are compared against the fixed Brewer MkIII at Sodankylä to detect any calibration drift during the campaign period.

# 3. Data Files

## 3.1  File location and naming

BTS data are not downloaded automatically. Daily CSV files must be placed manually into the local ./BTS_data/ directory. The expected filename pattern is:

YYYYMMDD_TOC_BTS_66639_V1.csv

One file is expected per measurement day.

## 3.2  File structure

Each CSV file contains a 15-line instrument metadata header followed by a column header row and data rows. Example:

Time (ISO 8601, GMT),Airmass,Ozone (DU),Ozone_STD,Ozone_Uncertainty

2026-04-15T05:45:45Z,3.165,381.06,1.528,5.549

| Column | Unit | Description |
| --- | --- | --- |
| Time | ISO 8601 (UTC) | Timestamp of each measurement (format: YYYY-MM-DDTHH:MM:SSZ) |
| Airmass | — | Optical air mass factor (= 1/cos(SZA)); unitless |
| Ozone (DU) | DU | Total ozone column in Dobson Units |
| Ozone_STD | DU | Standard deviation of ozone over the averaging window |
| Ozone_Uncertainty | DU | Combined measurement uncertainty (instrument + statistical) |

Values are already in Dobson Units — no unit conversion is required in the processing scripts.

## 3.3  How data are read by comparison scripts

The comparison scripts (gs_comparison.py, compare.py) scan ./BTS_data/ for all *.csv files matching the requested date range. For each file, the script:

- Skips the 15-line metadata header:

- Locates the column named: "Ozone (DU)" in the header row (dynamic column detection, not fixed index)

- Parses each subsequent data row: extracting Time and Ozone (DU) values

- Filters rows by date range: and assembles the time series for comparison with other instruments

No unit conversion is applied — BTS ozone values are used directly in Dobson Units as provided in the CSV.

# 4. Scientific Context & Limitations

## 4.1  Strengths

- WMO-traceable reference: Calibrated against the WMO World Standard Group; provides an independent, absolute reference for total ozone at Sodankylä.

- High temporal resolution: Measurements every 2–5 minutes throughout the day, enabling intra-day variability studies.

- Cross-validation capability: As a transportable instrument, the BTS can be directly compared side-by-side with the fixed Brewer MkIII to detect calibration biases in either instrument.

- Simple CSV format: Data are stored in a straightforward, human-readable format with no proprietary encoding; easy to integrate into Python/R analysis workflows.

## 4.2  Limitations

- Manual data transfer: No automated download. Files must be manually retrieved from the instrument and placed in BTS_data/. This can cause gaps if the transfer is delayed.

- Direct-sun only: Like all Brewer instruments, the BTS requires a clear view of the solar disk. Cloud cover causes data gaps.

- No vertical profile: Total column only. BTS data must be combined with ECC ozonesonde or DIAL lidar profiles for altitude-resolved information.

- Polar night gap: No measurements when the sun is below the horizon.

- Not to be confused with BTS DIAL Lidar: Both instruments share the acronym “BTS” at Sodankylä. The DIAL lidar measures vertical ozone profiles by laser; the BTS #66639 measures total column by UV photometry. File paths and formats are entirely different.

FMI Sodankylä Arctic Research Station  ·  BTS Brewer Transportable Spectrometer #66639 — Technical Documentation  ·  June 2026