TECHNICAL DOCUMENTATION

Pandora #180

Direct-Sun & Sky Spectrometer — SciGlob / NASA

NO₂ · O₃ · HCHO · H₂O · Tropospheric columns

FMI — Sodankylä Arctic Research Station  ·  67°22’N, 26°38’E

# 1. Instrument Overview

The Pandora Spectrometer System is a compact, fibre-fed UV–visible spectrometer designed by NASA Goddard Space Flight Center (GSFC) and commercialized by SciGlob Instruments & Services. It was developed specifically to support satellite validation campaigns — in particular for the OMI, TROPOMI (Sentinel-5P), and TEMPO instruments. The Pandora installed at Sodankylä is unit #180 and is part of the global Pandonia Global Network (PGN), a joint initiative of NASA and ESA to provide standardized reference observations for satellite trace gas validation at ground level.

Unlike the Brewer (which measures only 5 discrete wavelengths) or the SAOZ (which measures only at twilight), Pandora acquires full UV–visible spectra (280–530 nm) at high cadence throughout the day, in both direct-sun and sky modes. This enables the retrieval of multiple trace gases simultaneously and provides the intra-day temporal resolution needed to track rapidly evolving atmospheric composition.

Key specifications Spectral range: 280–530 nm (UV–visible) Spectral resolution: ~0.5 nm (FWHM) Detector: Back-illuminated CCD, 2048 pixels, thermoelectrically cooled Sun tracker: 2-axis motorized tracker; autonomous solar tracking Measurement modes: Direct sun (DS), zenith sky (ZS), off-axis (multi-angle) Temporal resolution: ~1–3 min (direct sun), continuous daylight Derived products: NO₂, O₃, HCHO, H₂O, O₄, BrO total & tropospheric columns Accuracy (NO₂): ~5–10% total column; ~10–15% tropospheric Accuracy (O₃): ~1–3% total column Network: Pandonia Global Network (PGN) — NASA/ESA

# 2. Measurement Principle

## 2.1  Direct-sun DOAS

Pandora points its tracking head directly at the solar disk and collects a continuous stream of spectra. The retrieval method is Direct-Sun DOAS: the measured spectrum is divided by a high-sun Fraunhofer reference spectrum (taken when the sun is high and the trace gas column is minimal) to produce an optical depth spectrum. The differential absorption structures of each target gas are then fitted simultaneously using their laboratory cross-sections. The result is a total slant column which is converted to a vertical column using an air mass factor (AMF) computed from the SZA and a climatological a priori profile.

## 2.2  Tropospheric column separation

By combining Pandora observations at multiple solar zenith angles (early morning vs. afternoon), it is possible to partially separate stratospheric and tropospheric contributions to the NO₂ column. The stratospheric column — well-mixed and slowly varying — is estimated using a climatological model or SAOZ measurements at twilight, and subtracted to yield the tropospheric NO₂. This tropospheric signal is the primary target for Sentinel-5P / TROPOMI validation at Sodankylä.

## 2.3  Multi-axis sky mode

In off-axis mode, Pandora scans different elevation angles to retrieve information on the vertical distribution of absorbers in the boundary layer using the MAX-DOAS (Multi-AXis DOAS) technique. This is particularly useful for characterizing the shape of the tropospheric NO₂ profile needed for satellite AMF corrections.

# 3. Data Output — File Format

## 3.1  Derived products

| Parameter | Unit | Mode | Notes |
| --- | --- | --- | --- |
| NO₂ total column | 10¹⁵ mol/cm² | DS / ZS | Primary validation product for TROPOMI |
| NO₂ tropospheric column | 10¹⁵ mol/cm² | DS | Stratosphere subtracted |
| O₃ total column | DU | DS | Secondary product |
| HCHO total column | 10¹⁵ mol/cm² | DS | Formaldehyde |
| H₂O total column | g/cm² | DS | Water vapour |
| O₄ slant column | 10⁴³ mol²/cm⁵ | DS | Oxygen dimer; cloud/aerosol proxy |
| BrO total column | 10¹³ mol/cm² | ZS | Reactive bromine |

## 3.2  PGN data format

Data are distributed by the Pandonia Global Network (pandonia-global-network.org) in ASCII columnar files, one per day. Files follow the PGN naming convention and include:

- Level 0: raw spectra (photon counts vs. wavelength per timestamp)

- Level 1.5: calibrated spectra with quality flags and solar zenith angle

- Level 2.0: retrieved trace gas columns with uncertainties and AMF

| Column (L2) | Description |
| --- | --- |
| UTC datetime | Timestamp of measurement |
| SZA | Solar zenith angle (°) |
| NO2_total | NO₂ total vertical column (10¹⁵ mol/cm²) |
| NO2_tropo | NO₂ tropospheric column |
| dNO2 | 1σ uncertainty |
| O3_total | O₃ vertical column (DU) |
| HCHO_total | Formaldehyde column |
| cloud_flag | 0 = clear, 1 = cloudy (based on O₄ threshold) |
| quality_flag | PGN quality flag (0 = best quality) |

## 3.3  Using the download script (download_pandora.py)

A Python download script is provided to retrieve Pandora Level 2 data directly from the PGN API. Configuration is done at the top of the file:

| Parameter | Value | Description |
| --- | --- | --- |
| SITE | "Sodankyla" | PGN site name |
| PAN_ID | 309 | Sodankylä instrument ID |
| SPECTROMETER | "1" | Spectrometer unit number |
| LEVEL | "L2" | Processing level |
| CODE | "rout2" | Product code (ozone total column) |
| DATE_START / DATE_END | "YYYY-MM-DD" | ISO date range |

The script iterates over each day in the requested range and queries the PGN file API:

GET /v1/files/{site}/{pan_id}/{spectrometer}/{level}?start=...&end=...&code=...

Downloaded files are saved as plain text into ./pandora_data/, using the naming convention:

pandonia_Sodankyla_{date}_L2_rout2.txt

Each file contains a 75-line header followed by tab-separated data rows. The key field for ozone is at column index 39 (0-based), labelled O3_total, in units of mol/m². Conversion to Dobson Units is done in the comparison scripts by multiplying by 2241:

O3_DU = O3_total [mol/m²] × 2241

# 4. Scientific Context & Limitations

## 4.1  Strengths

- Sentinel-5P / TROPOMI validation: Pandora is the reference instrument of choice for TROPOMI NO₂ validation. Its direct-sun geometry minimises the AMF uncertainty that affects zenith instruments.

- Multi-species, high cadence: Six trace gases from a single instrument, every 1–3 minutes, all day long. Essential for capturing pollution events and photochemical cycles.

- Global network standardisation: PGN provides harmonised L1–L2 processing for all Pandora units worldwide, enabling systematic cross-station comparisons.

- Compact and autonomous: Self-tracking, no cryogenic cooling, low maintenance. Suitable for long-term unattended operation.

## 4.2  Limitations

- Clear sky required: Direct-sun mode fails under cloudy conditions. The O₄ column is used as a proxy to flag cloud-contaminated measurements, but thick overcast periods create data gaps.

- AMF uncertainty for tropospheric column: The separation of stratospheric and tropospheric NO₂ introduces an additional ~10–20% uncertainty, dependent on the a priori profile and AMF model.

- No vertical profile: Total and tropospheric columns only; altitude-resolved information requires MAX-DOAS or ozonesonde data.

- Temperature sensitivity: The CCD is thermoelectrically cooled but not perfectly stable; temperature fluctuations affect dark current and spectral calibration.

FMI Sodankylä Arctic Research Station  ·  Pandora #180 — Technical Documentation  ·  June 2026